## Slide 1
The Method of passive SAMPLING for measuring mercury concentration in the atmosphere

## Slide 2: Standard Sampling of Mercury in The Atmosphere (Active Sampling Method)
Air
Sampling pump
Sampling tube
Active sampling
Flow rate 0.5L/min

## Slide 3
・No Power Supply
Compact and Cheap
Can Study for Long Term Contamination Level
Wide Range and Multi-point Measurement
Suitable for Personal , Indoor and Field Monitoring
Disk type
Tube type
Advantage of Passive Sampling

## Slide 4
Usage of Passive Sampling
VOC monitoring
Measurement of parsonal exposure

## Slide 5: Principle of Passive Sampling
Gaseous Hg（Target）
Diffusive Surface
Adsorbing Surface
The target substance permeation the tube and diffuses to the adsorbent surface in the tube and is collected
The amount of collection depends on the sampling time
Electronic microscope image
Tube fo passive sampler

## Slide 6: Calculation of Atmospheric Concentration
C：Atmospheric concentration of mercury（ng/m3)
Ws : Amount of adsorbed mercury (ng)
Wb : Amount of sample blank (ng)
SR : Sampling rate (ng/(ng/m3)/day)
T : Sampling time (day)
Before passive sampler monitoring, sampling rate must be determined.
C＝（Ws-Wb）/（SR×T）

## Slide 7: Passive Sampling vs Active Sampling
During the same period, passive and active sampling will be performed concurrently.
The amount of mercury collected for each is measured, and the SR is calculated from the correlation.
C＝（Ws-Wb）/（SR×T）
Ws
SR=
(C×T）
Ws (ng)
C (ng/m3)
VS
(ng/(ng/m3)/day)

## Slide 8: Verification Test for SR Determine
Sampling period (day) | 7
Active sampling volume (m3) | 0.998
Amount of mercury (Active) (ng) | 1.43
Atmospheric concentration of mercury （ng/m3) | 1.43
Amount of mercury (Passive) (ng) | 0.727
Sampling rate | 0.073

## Slide 9: Now let's move on to the practical training !